## Background Information
Audit of GFR Measurement - 2026
Nuclear Medicine Software Quality Group
1. Sampling method and counting
You have been supplied with 16 sets of real clinical data (clinical details included where available) obtained after IV administration of Tc‑99m DTPA. The adult DRL was 10 MBq and was scaled for paediatric patients according to published guidance. Blood samples were taken after administration, and equal 1 ml aliquots of plasma were prepared for counting. Each plasma sample was counted for 5 minutes, alongside an equal (1 ml) volume of a diluted standard, plus an empty tube to provide a background reading. The sample counter has already applied decay correction for the duration of the count. Background counts have been subtracted from the patient samples, and the results are presented in counts per minute (cpm). If your GFR calculator requires a “background count rate,” please enter 0.
Please use the supplied test data and your knowledge of your local GFR procedure to complete the Main Audit Results Section and Supplementary Section, and return the completed document to Laura Perry (laura.perry2@nhs.net).
The submitted datasets will be collated to compare GFR results across participating sites. You will receive feedback showing how your centre’s performance compares to all submitted datasets.
2. Adapting your normal method
2.1 Sample counting times
The sample counter has already applied decay correction for these studies. If your calculator asks for times of sample measurements (hh:mm), you may enter the same time for all samples (e.g. 14:00), as no further decay correction is required for this dataset.
2.2 Volume of dilution and sample volumes
A problem may arise if your local method involves a volume of dilution for your standard which is different to 1000ml and/or if the sample volumes used for the standard and plasma samples are different from 1ml. If your GFR calculator allows manual input of these issues then please use the values provided (1000ml, 1ml, 1ml respectively). If these values are hard‑coded into your calculator, you may need to adjust the supplied test data; typically only the standard count (Cs) needs to be recalculated. This is best illustrated by way of an example:
If your local method assumes:
Volume of dilution for the standard = 250ml (Vds)
Standard sample volume = 2ml (Vs)
Plasma sample volume = 2ml (Vp)
and that these volumes are "assumed" as they are hard-coded into your GFR calculator.
The value of Cs to be used with your GFR calculator should be calculated as follows:
where Cs = counts from the standard to be used with your program (cpm)
Cg = counts from the standard given in the test data (cpm)
Vds = volume of dilution of the standard assumed by your program (ml)
In the example given above, if the standard sample count rate given in the test data (Cg) is 220200, the value to be entered (Cs) would be 880800 cpm.
2.3 Use of activity rather than weight for calibration purposes
Some GFR calculators assume standard and patient activities are measured using a radionuclide calibrator rather than derived from syringe weights. Such GFR calculators may ask for activity in MBq (before and after injection) for both the standard and patient injections. You may be able to substitute the weights given in the test data for activity because it is the ratio of the activity measurment that is important. Use the following substitution:
Patient activity (MBq) before administration = Weight of the full patient syringe (g)
Patient activity (MBq) after administration = Weight of the residual patient syringe (g)
Standard activity (MBq) before injection = Weight of the full standard syringe (g)
Standard activity (MBq) after injection = Weight of the residual standard syringe (g)
2.4 Use of different blood sample timings
The test data provided covers the typical range over which blood sampling is obtained as set out in the BNMS GFR guidelines 2018 and understood to be commonly used. However, the timing of the blood samples may be different to your routine GFR measurement method. If your analysis method uses fewer blood samples, please select the samples that most closely match the times at which you would normally take them. If your method requires a sample at particular times, you could plot the data and interpolate. Whatever method you use, please indicate it clearly on the Results Sheet under "Additional Comments" for each patient.
2.5 Use of a sensitivity factor rather than standard
If your GFR calculator uses a counter sensitivity factor instead of a standard made from a sample of the patient radiopharmaceutical vial a counter calibration factor has been provided (cpm/kBq in 1ml) and the administered activity at time of sample counting (MBq) has been provided for use. Please add a note to highlight that you have used this method.
2.6 Your GFR counter applies a decay correction for the sample counting duration.
The data assumes decay correction was applied by the sample counter. If your GFR calculator also applies decay correction, please ensure correction is not applied twice. You could enter the time of sample counting as the same value for each patient and standard sample. Alternatively you could "uncorrect" the provided count data for the sample counting durations that your calculator expects - the calculator can then reapply the same magnitude correction and the GFR will be calculated from the supplied sample counts.
3. Expressing GFR
Please report GFR in ml/min and ml/min/1.73m2 (ie normalised for body surface area) if your software outputs both values. Use the final result(s) after all QC checks or corrections, this should match the result your centre would normally report.
Finally a few points about size normalisation:
1. If you do not normalise to body surface area (BSA), clearly state this and leave those sections blank.
2. If you do normalise to BSA, include the formula used and a reference if possible.
3. If you use a different normalisation parameter, replace “ml/min/1.73m²” with the correct unit, and explain the method used.
Many thanks for participating in this audit. Laura Perry Imperial College Healthcare NHS Trust, on behalf of the IPEM Nuclear Medicine Software Quality Group
Please also direct questions or comments to the email address below
laura.perry2@nhs.net

## Audit Data
Audit of GFR Measurement - 2026
Nuclear Medicine Software Quality Group
This sheet contains plasma count rate data from 16 patients following injection of approx Tc99m DTPA
Note: See the information in the Background Information sheet for details of how the samples were obtained, counted, etc, plus advice on whether you need to modify any of the values before entering them into your particular program.
In summary:
· All counts given are background subtracted.
· The standard and plasma samples contained the same volume (1ml).
· Samples were counted for 5 minutes, and results expressed in cpm.
Please record your results in the "Main Audit Result Section" worksheet
Patient details
Patient i.d | Patient1 | Patient2 | Patient3 | Patient4 | Patient5 | Patient6 | Patient7 | Patient8 | Patient9 | Patient10 | Patient11 | Patient12 | Patient13 | Patient14 | Patient15 | Patient16
Clinical Indications | GFR assessment prior to treatment for adenosquamous lung malignancy. | Unexplained low e-GFR | Reduced renal function post kidney donation. | Potential kidney donor | Potential kidney donor | Potential kidney donor | Potential kidney donor | GFR assessment after induction chemotherapy. | Previous chemotherapy, assessment prior to next chemotherapy. | GFR assessment prior to carboplatin chemotherapy | Ovarian cancer. For adjuvant chemo. | Low eGFR but well muscled and need true GFR | Advanced ovarian cancer. Ascites and pleural effusion. For carboplatin chemo to start ASAP | ovarian cancer - for primary chemo - carboplatin. Known sympotomatic ascites | pre chemotherapy for carboplatin. Known pleural effusion drained 3 weeks ago. | Metastatic NET for Lutathera - requires baseline GFR. Patient has right pleural effusion
Most recent eGFR (where available) ml/min | 48 | 58 | 48 | 72 | 89 | 83 | > 90 | Unknown | Unknown | 30 | 12 | 24 | Unknown | Unknown | Unknown | Unknown
Age (yrs) | 79 | 42 | 39 | 61 | 64 | 41 | 39 | 1 | 6 | 72 | 49 | 57 | 73 | 68 | 43 | 57
Sex | Male | Male | Female | Female | Male | Male | Male | Male | Female | Male | Female | Male | Female | Female | Female | Female
Height (m) | 1.71 | 1.76 | 1.74 | 1.71 | 1.775 | 1.72 | 1.76 | 0.795 | 1.23 | 1.81 | 1.68 | 1.7 | 1.61 | 1.476 | 1.57 | 1.61
Weight (kg) | 77 | 69.3 | 78 | 90 | 67 | 81 | 100 | 11.36 | 25.5 | 85 | 145 | 75 | 52 | 44.1 | 50.9 | 63.8
Tc-99m DTPA standard
Weight of syringe + patient injection (g) | 4.6529 | 4.6581 | 4.491 | 5.0479 | 4.9173 | 4.446 | 4.5338 | 4.1057 | 4.289 | 4.5033 | 4.3351 | 4.059 | 4.3597 | 4.65 | 4.59 | 4.423
Weight of empty syringe (g) | 3.8599 | 3.8861 | 3.8753 | 3.8844 | 3.9276 | 3.9034 | 3.9033 | 3.8681 | 3.8509 | 3.8962 | 3.7902 | 3.5022 | 4.0469 | 3.431 | 3.432 | 3.438
Weight of syringe + standard (g) | 4.3992 | 4.3007 | 4.2416 | 4.3139 | 4.374 | 4.3775 | 4.3456 | 4.2821 | 4.3374 | 4.9054 | 4.6719 | 4.5484 | 4.4268 | 4.5 | 4.444 | 4.339
Weight of empty syringe (g) | 3.8744 | 3.8774 | 3.8112 | 3.8772 | 3.9008 | 3.9455 | 3.8462 | 3.8994 | 3.8393 | 3.8712 | 3.7652 | 3.4772 | 4.0959 | 3.462 | 3.475 | 3.419
Dilution of standard (ml) | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000
For centres using a Standard
Standard Count rate (cpm) | 115204 | 102861 | 101955 | 66416 | 96931 | 113156 | 125701 | 87334 | 132817 | 98830 | 99952 | 99726 | 69400 | 99952 | 106373 | 105245
For centres using a counter calibration factor:
Calibration factor (cpm/kBq in 1ml sample) | 44000 | 44000 | 44000 | 44000 | 44000 | 44000 | 44000 | 44000 | 44000 | 44000 | 44000 | 44000 | 44000 | 44000 | 44000 | 44000
Patient administered activity at time of counting / MBq | 3.9563 | 4.2635 | 3.3147 | 4.0216 | 4.6075 | 3.2301 | 3.6068 | 1.2323 | 2.6549 | 1.3184 | 1.3652 | 1.1781 | 1.4909 | 2.6677 | 2.8891 | 2.5609
Plasma samples
Time after injection for 1st sample (mins) | 125 | 120 | 133 | 120 | 121 | 122 | 117 | 120 | 120 | 205 | 185 | 127 | 125 | 120 | 115 | 124
Time after injection for 2nd sample (mins) | 184 | 180 | 180 | 180 | 183 | 180 | 179 | 182 | 185 | 1541 | 1436 | 188 | 176 | 240 | 236 | 235
Time after injection for 3rd sample (mins) | 241 | 241 | 240 | 240 | 245 | 240 | 240 | 243 | 233 |  |  | 242 | 235 | 368 | 355 | 364
Time after injection for 4th sample (mins) | 300 | 301 | 300 | 300 | 317 | 300 | 297 | 297 | 300 |  |  | 1460 | 350 | 1415 | 1354 | 1383
Plasma counts (cpm) for 1st sample | 7474 | 8102 | 4897 | 4808 | 5609 | 3300 | 3242 | 5630 | 7130 | 2164 | 2067 | 2813 | 2833 | 6667 | 3969 | 3685
Plasma counts (cpm) for 2nd sample | 6395 | 6188 | 3682 | 3626 | 3939 | 2140 | 1960 | 3229 | 4179 | 280 | 646 | 2603 | 2415 | 4345 | 1782 | 2443
Plasma counts (cpm) for 3rd sample | 5677 | 4714 | 2739 | 2531 | 2780 | 1405 | 1264 | 1949 | 2939 |  |  | 2423 | 1982 | 2943 | 908 | 1445
Plasma counts (cpm) for 4th sample | 4917 | 3639 | 2027 | 1914 | 1846 | 965 | 896 | 1236 | 1758 |  |  | 814 | 1435 | 353 | 78 | 139

## Main Audit Result Section
|  |  |  | Nuclear Medicine Software Quality Group | [9 empty] | Yes
No
|  |  |  | https://www.ipem.ac.uk/about/organisational-structure/special-interest-groups/nuclear-medicine-special-interest-group/nuclear-medicine-software-group-latest-news-and-current-audit/ | [9 empty] | Don't know
Audit of GFR Measurement - 2026
Please fill in the green cells with your answers. Try and avoid leaving questions unanswered as much as possible.
1 | Hospital or Trust
2 | Address
3 | NHS Region
4 | Contact Person
5 | Email
6 | Telephone
7 | Results | Patient # | GFR (ml/min) | Normalised GFR (ml/min/1.73m^2) | ECFV (litres) | Intercept | halftime (min) | slope | Which samples were used in the calculation? | Which method was used to calculated this GFR? |  | Comment on the quality of the fits or note action undertaking in processing the study as would be per normal practise
Patient 1
Patient 2
Patient 3
Patient 4
Patient 5
Patient 6
Patient 7
Patient 8
Patient 9
Patient 10
Patient 11
Patient 12
Patient 13
Patient 14
Patient 15
Patient 16
Comments
Please state other means of normalising GFR if not by BSA (with a reference)?
Please provide the units for slope and intercept (if included)
13 | Please complete the following table to describe the GFR calculation methods in use at your centre. Where multiple calculation methods are in use, please describe each together with a description of when they are used.
|  | GFR calculation method: |  | Is this method in use at your centre? ….. Yes / No / Don't know | What time are the blood sample(s) taken? |  | How is the time of the blood sample(s) determined? |  | Do you apply correction(s) in the calculation to account for the missing early exponential AUC (using 2-4 hour sampling only) |  | If yes, which correction(s) are used? |  | Published reference(s) for these correction(s)?
Slope-intercept
Single sample
Slope only
Area Under Curve (eg Wickham et al)
14 | Do you routinely apply any other correction factors? If so please list them, preferably with a reference:
|  | # 1 | Correction applied for |  | Ref 1
|  | # 2 | Correction applied for |  | Ref 2
Comment (if needed):
15 | Which of the following Quality Control methods are used at your site? Please add details about how the checks are implemented.
|  | Quality Control method: |  | Is this method in use at your centre? ….. Yes / No/ Don't know | Further details about how the QC check is performed
Check for third space eg oedema, ascites or effusion
Compromised blood counts due to recent nuclear medicine or radionuclide therapy
Inaccuracy in measurement or composition of standard
Incorrect measurement of administered activity
Check that all activity is administered intravenously
Inaccuracy of blood sample times
Inaccuracy of pipetted volumes
Consistency of standard and patient sample volumes
Counts recorded erroneously from other sources
Haemolysis of plasma samples
Check that count data from correct patient is used
Transcription errors
16 | Analysis method (please place an 'x' in the appropriate following box and fill in the langauge or software box if required)
Manual (log-lin graph paper and calculator)
|  |  | Computer program written in |  |  | (language)
|  |  | Computer spreadsheet developed using |  |  | (name of software)
17 | If software based, please place an 'x' in the appropriate following box
Developed in-house
Obtained from another department
|  |  | Purchased from a commercial source. |  | Name of supplier / manufacturer
18 | What BSA formula do you use, if any?
Adults - Male
Reference(s) for above:
Adults - Female
Reference(s) for above:
Children and infants
Reference(s) for above:
Do you normalise GFR to BSA before applying the slope-intercept AUC correction?

## Supplementary Section
Supplementary Section
19 |  | How long has the above method/program been in use locally? | [5 empty] | yrs (approx)
20 |  | Which radiopharmaceutical do you routinely use for GFR?
Yes
|  |  | If DTPA is used, please answer the following questions: | [5 empty] | No
Don't know
What preparation is used?
Do you sub-dispense the DTPA?
Have you / do you perform any tests or audits to evaluate protein binding?
21 |  | How many GFR studies do you perform each year ? | [5 empty] | (approx)
|  |  | Of these patients, what percent would you estimate are |  |  |  |  | Renal patients
Renal function assesement in living donors
Dermatology
Oncology patients
Other
22 |  | Please state and describe whether you would do anything differently with a request for a GFR measurement in a patient in the following situation:
|  |  | Patient circumstances: |  |  | Would you do anything differently? | Please describe your practice for this patient group
A patient who had a known fluid collection (oedema, ascites, pleural effusions etc):
A patient who had limb(s) amputation:
A patient who has recently had an IV administration of CT contrast
A patient who has recently had an IV administration of MR contrast
A patient with a dual lumen central line
23 |  | Do you use a standard when performing GFR studies? ….. Yes / No / Don't know
If yes, what volume of standard dilution do you use?
What volume of the standard samples do you use?
If no standard is used, please describe your practice:
24 |  | What volume of plasma samples do you use?
Do you apply any corrections for deadtime?
Do you have any restrictions on the maximum delay between taking a sample and counting it which would make the study unreportable?
If yes, how is this maximum determined?
25 |  | Please list any other measurements you report alongside GFR and normalised GFR:
26 |  | What are your normal ranges for GFR measurements in
|  | Adults - Male |  |  |  |  | (include units)
|  | Adults - Female |  |  |  |  | (include units)
Please state if you use different normal ranges for different age ranges
Reference(s) for above:
|  | Children and infants |  |  |  |  | (include units)
Reference(s) for above:
If you previously used Cr-51 EDTA and now use Tc-99m DTPA:
Do you use the same normal range values?
27 |  | Did you take part in the previous GFR audit in 2013?
Did you change your method since then?
28 |  | Additional comments
THE END. Thank you for your time.
Please return a completed version of this audit spreadsheet to the contact details below:
Please also direct questions or comments to the email address below
Contact Details
Laura Perry |  | laura.perry2@nhs.net |  |  | Imperial College Healthcare NHS Trust